## Slide 1: BIOMEDICAL INSTRUMENTATION
MODULE-2
PROF. DR. JOYANTA KUMAR ROY
DEPARTMENT OF APPLIED ELECTRONICS & INSTRUMENTATION ENGINEERING
NARULA INSTITUTE OF TECHNOLOGY
WWW.dr-joyanta-kumar–roy.com

## Slide 2: Module - 2 Fundamentals of Biomedical Instrumentation

## Slide 3
General Instrumentation Knowledge
Basic Concepts of Medical Instrumentation
Origin of Bio potentials
Bio potential Electrodes
Effects of Interface Impedances
Amplifiers
Operational Amplifiers
(Basic circuit configurations, parameters, data sheets)
Instrumentation Amplifiers
Bio potential Amplifiers

## Slide 5: Medical and Physiological Parameters #1

## Slide 6: Medical and Physiological Parameters #2

## Slide 7: Medical and Physiological Parameters #3

## Slide 12: Generalized Instrumentation System
Perceptible
output
Output
display
Control
And
feedback
Signal
processing
Data
transmission
Data
storage
Variable
Conversion
element
Sensor
Primary
Sensing
element
Measurand
Calibration
signal
Radiation,
electric current,
or other applied
energy
Power
source
The sensor converts energy or information from the measurand to another form (usually electric). This signal is processed and displayed so that humans can perceive the information. Elements and connections shown by dashed lines are optional for some applications.

## Slide 15: Generalized Static Characteristics of Instrumentation System
Accuracy
  Difference between the true value and the measured value divided by the true value
  True value is a traceable standard to NIST
  Expressed as a percentage of the reading or of full scale, or ± the smallest analog division
Precision
  Number of distinguishable alternatives (eg., a meter with a five digit readout)
Resolution
  Smallest incremental quantity that can be measured with certainty
Reproducibility
  Produces the same output for a given input over a period of time
Statistical Control
  Systematic and bias error can be removed by calibration
  Random errors can be removed by taking multiple measurements and averaging the results

## Slide 16: Simplified Electrocardiographic Recording System
Electrodes
60-Hz
ac magnetic
field
Displacement
currents
Differential
amplifier
+
-
+
Vcc
-Vcc
Z1
Zbody
Z2
vo
vecg
Two possible interfering inputs are stray magnetic fields and capacitively coupled noise. Orientation of patient cables and changes in electrode-skin impedance are two possible modifying inputs. Z1 and Z2 represent the electrode-skin interface impedances.

## Slide 17: Basic Biopotential Amplifier Requirements
Purpose:
  To provide voltage and/or current gain to increase the amplitude of weak electric signals of biological origin
Features:
  High input impedance (minimize the loading effects of the amplifier inputs)
  Protection circuitry (limit the possibility of introducing dangerous microshocks or macroshocks at the input terminals of the amplifier)
  Low output impedance (low with respect to the load being driven)
  Adequate output current (to supply the power needed to drive the load)
  Bandlimited frequency response (match the frequency response of the signal being measured to eliminate out-of-band noise)
  Quick calibration (include a signal source and a number of selectable fixed gains settings)
  High common-mode rejection for differential amplifiers (common mode signals are frequently larger than the biopotentials being measured)
Additional specific requirements for each application

## Slide 18: Biopotential Electrodes by Dr. Joyanta Kumar Roy

## Slide 19: Biopotential Electrodes Outline
The Electrode-Electrolyte Interface
Polarization
Polarizable and Nonpolarizable Electrodes
Electrode Behavior & Circuit Models
The Electrode-Skin Interface & Motion Artifact
Body-Surface Recording Electrodes
Internal Electrodes
Electrode Arrays
Microelectrodes
Electrodes for Electric Stimulation of Tissue
Practical Hints in Using Electrodes

## Slide 20: Biopotential Electrodes – The Basics
The interface between the body and electronic measuring devices
Conduct current across the interface
Current is carried in the body by ions
Current is carried in electronics by electrons
Electrodes must change ionic current into electronic current
This is all mediated at what is called the Electrode-Electrolyte Interface or the Electrode-Tissue Interface

## Slide 24: Current Flow at the Electrode-Electrolyte Interface
Figure 5.1 The current crosses it from left to right. The electrode consists of metallic atoms C. The electrolyte is an aqueous solution containing cations of the electrode metal C+ and anions A-.
Electrons move in opposite direction to current flow
Cations (C+ ) move in same direction as current flow
Anions (A– ) move in opposite direction of current flow
Chemical oxidation (current flow right) - reduction (current flow left) reactions at the interface:
    C C + + e – (5.1)
    A – A + e – (5.2)
No current at equilibrium
 Electron flow
 Ion- flow
+ Current flow 
Ion+ flow 

## Slide 25: Half-Cell Potential
When metal (C) contacts electrolyte, oxidation (C  C + + e –) or reduction (A-  A + e –) begins immediately.
Local concentration of cations at the surface changes.
Charge builds up in the regions.
Electrolyte surrounding the metal assumes a different electric potential from the rest of the solution.
This potential difference is called the half-cell potential ( E0 ).
Separation of charge at the electrode-electrolyte interface results in a electric double layer (bilayer).
Measuring the half-cell potential requires the use of a second reference electrode.
By convention, the hydrogen electrode is chosen as the reference.
J.K.Roy
25

## Slide 26: Half-Cell Potentials of Common Metals at 25 ºC
Metal Potential E0 (volts)
Al - 1.706
Zn - 0.763
Cr - 0.744
Fe - 0.409
Cd - 0.401
Ni - 0.230
Pb - 0.126
H 0.000
AgCl + 0.223
Hg2Cl2 + 0.268
Cu + 0.522
Ag + 0.799
Au + 1.680
By definition: Hydrogen is bubbled over a platinum electrode and the potential is defined as zero.

## Slide 27: Electrode Polarization
Standard half-cell potential ( E0 ):
  Normally E0 is an equilibrium value and assumes zero-current across the interface.
  When current flows, the half-cell potential, E0 , changes.
Overpotential ( Vp ):
  Difference between non-zero current and zero-current half-cell potentials; also called the polarization potential (Vp).
Components of the overpotential ( Vp ):
  Ohmic ( Vr ): Due to the resistance of the electrolyte (voltage drop along the path of ionic flow).
  Concentration ( Vc ): Due to a redistribution of the ions in the vicinity of the electrode-electrolyte interface (concentration changes).
  Activation ( Va ): Due to metal ions going into solution (must overcome an energy barrier, the activation energy) or due to metal plating out of solution onto the electrode (a second activation energy).
Vp = Vr + Vc + Va (5.4)

## Slide 28: Nernst Equation
Governs the half-cell potential:
where
E – half-cell potential
E0– standard half-cell potential
(the electrode in an electrolyte with unity
activity at standard temperature)
R – universal gas constant [ 8.31 J/(mol K) ]
T – absolute temperature in K
n – valence of the electrode material
F – Faraday constant [ 96,500 C/(mol/valence) ]
– ionic activity of cation Cn+
(its availability to enter into a reaction)
(5.6)

## Slide 29: Hydrogen Electrode
Standard Hydrogen Electrode (SHE)
The SHE is the universal reference for reporting relative half-cell potentials. It is a type of gas electrode and was widely used in early studies as a reference electrode, and as an indicator electrode for the determination of pH values. The SHE could be used as either an anode or cathode depending upon the nature of the half-cell it is used with. The SHE consists of a platinum electrode immersed in a solution with a hydrogen ion concentration of 1.00M. The platinum electrode is made of a small square of platinum foil which is Platonized (known as platinum black). Hydrogen gas, at a pressure of 1 atmosphere, is bubbled around the platinum electrode. The platinum black serves as a large surface area for the reaction to take place, and the stream of hydrogen keeps the solution saturated at the electrode site with respect to the gas. It is interesting to note that even though the SHE is the universal reference standard, it exists only as a theoretical electrode which scientists use as the definition of an arbitrary reference electrode with a half-cell potential of 0.00 volts. (Because half-cell potentials cannot be measured, this is the perfect electrode to allow scientists to perform theoretical research calculations.) The reason this electrode cannot be manufactured is due to the fact that no solution can be prepared that yields a hydrogen ion activity of 1.00M. hydrogen electrode is made by adding platinum black to platinum wire or a platinum plate. It is immersed in the test solution and an electric charge is applied to the solution and platinum black with hydrogen gas. The hydrogen-electrode method is a standard among the various methods for measuring pH. The values derived using other methods become trustworthy only when they match those measured using hydrogen electrode method. However, this method is not appropriate for daily use because of the effort and expense involved, with the inconvenience of handling hydrogen gas and great influence of highly oxidizing or reducing substances in the test solution.

## Slide 30: Polarizability & Electrodes
Perfectly polarizable electrodes:
  No charge crosses the electrode when current is applied
  Noble metals are closest (like platinum and gold); they are difficult to oxidize and dissolve.
  Current does not cross, but rather changes the concentration of ions at the interface.
  Behave like a capacitor.
Perfectly non-polarizable electrodes:
  All charge freely crosses the interface when current is applied.
  No overpotential is generated.
  Behave like a resistor.
  Silver/silver-chloride is a good non-polarizable electrode.

## Slide 31: The Classic Ag/AgCl Electrodes
Features:
  Practical electrode, easy to fabricate.
  Metal (Ag) electrode is coated with a layer of slightly soluble ionic compound of the metal and a suitable anion (Cl).
Reaction 1: silver oxidizes at the Ag/AgCl interface
    Ag Ag + + e – (5.10)
Reaction 2: silver cations combine with chloride anions
Ag + + Cl – Ag Cl (5.11)
AgCl is only slightly soluble in water so most precipitates onto the electrode to form a surface coating.
Figure 5.2 A silver/silver chloride electrode, shown in cross section.

## Slide 32: Ag/AgCl Electrodes
Solubility product ( Ks ): The rate of precipitation and of returning to solution. At equilibrium:
      Ks = aAg+ x aCl - (5.12)
The equation for the half-cell potential becomes
E = E0Ag + ln ( Ks ) - ln ( aCl - )
      (5.15)
Determined by the activity of the chloride ion. In the body, the activity of Cl – is quite stable.
RT
nF
RT
nF
constant

## Slide 33: Ag/AgCl Fabrication
Electrolytic process
Large Ag/AgCl electrode serves as the cathode.
Smaller Ag electrode to be chloridized serves as the anode.
A 1.5 volt battery is the energy source.
A resistor limits the current.
A milliammeter measures the plating current.
Reaction has an initial surge of current.
When current approaches a steady state (about 10 µA), the process is terminated.
Cathode Anode
A
Electrochemical Cell

## Slide 34: Sintered Ag/Ag Electrode
Sintering Process
A mixture of Ag and AgCl powder is pressed into a pellet around a silver lead wire.
Baked at 400 ºC for several hours.
Known for great endurance (surface does not flake off as in the electrolytically generated electrodes).
Silver powder is added to increase conductivity since AgCl is not a good conductor.
Figure 5.3

## Slide 38: Calomel Electrode
Calomel is mercurous chloride (Hg2Cl2).
Approaches perfectly non-polarizing behavior
Used as a reference in pH measurements.
Calomel paste is loaded into a porous glass plug at the end of a glass tube.
Elemental Hg is placed on top with a lead wire.
Tube is inserted into a saturated KCl solution in a second glass tube.
A second porous glass plug forms a liquid-liquid interface with the analyte being measured.
Hg2Cl2
Hg
K Cl
Electrolyte being measured
porous glass plug
Lead Wire

## Slide 39: Electrode Circuit Model
Ehc is the half-cell potential
Cd is the capacitance of the electric double layer (polarizable electrode properties).
Rd is resistance to current flow across the electrode-electrolyte interface (non-polarizable electrode properties).
Rs is the series resistance associated with the conductivity of the electrolyte.
At high frequencies: Rs
At low frequencies: Rd + Rs
Figure 5.4

## Slide 40: Ag/AgCl Electrode Impedance
Figure 5.5 Impedance as a function of frequency for Ag electrodes coated with an electrolytically deposited AgCl layer. The electrode area is 0.25 cm2. Numbers attached to curves indicate the number of mAs for each deposit.

## Slide 41: Nichel- & Carbon-Loaded Silicone
Figure 5.6
Electrode area is 1.0 cm2

## Slide 42: Skin Anatomy
Figure 5.7

## Slide 43: Electrode-Skin Interface Model
Motion artifact:
Gel is disturbed, the charge distribution is perturbed changing the half-cell potentials at the electrode and skin.
Minimized by using non-polarizable electrode and mechanical abrasion of skin.
Skin regenerates in 24 hours.
Figure 5.8 A body-surface electrode is placed against skin, showing the total electrical equivalent circuit obtained in this situation. Each circuit element on the right is at approximately the same level at which the physical process that it represents would be in the left-hand diagram.
Sweat glands
and ducts
Electrode
Epidermis
Dermis and
subcutaneous layer
Ru
Re
Ese
Ehe
Rs
Rd
Cd
EP
RP
CP
Ce
Gel

## Slide 44: Metal Electrodes
Figure 5.9 Body-surface biopotential electrodes (a) Metal-plate electrode used for application to limbs. (b) Metal-disk electrode applied with surgical tape. (c) Disposable foam-pad electrodes, often used with electrocardiograph monitoring apparatus.

## Slide 45: Metal Suction Electrodes
A paste is introduced into the cup.
The electrodes are then suctioned into place.
Ten of these can be with the clinical electrocardiograph – limb and precordial (chest) electrodes
Figure 5.10

## Slide 46: Floating Metal Electrodes
Mechanical technique to reduce noise.
Isolates the electrode-electrolyte interface from motion artifacts.
Figure 5.11 (a) Recessed electrode with top-hat structure. (b) Cross-sectional view of the electrode in (a). (c) Cross-sectional view of a disposable recessed electrode of the same general structure shown in Figure 5.9(c). The recess in this electrode is formed from an open foam disk, saturated with electrolyte gel and placed over the metal electrode.
Double-sided
Adhesive-tape
ring
Insulating
package
Metal disk
Electrolyte gel
in recess
(a)
(b)
(c)
Snap coated with Ag-AgCl
External snap
Plastic cup
Tack
Plastic disk
Foam pad
Capillary loops
Dead cellular material
Germinating layer
Gel-coated sponge

## Slide 47: Flexible Body-Surface Electrodes
Carbon-filled silicone rubber
(b) Flexible Mylar film with Ag/AgCl electrode
(c) Cross section of the Mylar electrode
Figure 5.12

## Slide 48: Percutaneous Electrodes
Insulated needle
(b) Coaxial needle
(c) Bipolar coaxial needle
(d) Fine wire, ready for insertion
(e) Fine wire, after insertion
(f) Coiled fine wire, after insertion
Figure 5.13

## Slide 49: Fetal Intracutaneous Electrodes
Figure 5.14
Suction needle electrode
Suction electrode (in place)
Helical electrode (attached by corkscrew action)

## Slide 50: Implantable Electrodes
Figure 5.15
Multielement depth electrode array
Wire-loop electrode
Cortical surface potential electrode

## Slide 51: Microfabricated Electrode Arrays
Figure 5.16 (a) One-dimensional plunge electrode array (b) Two-dimensional array, and (c) Three-dimensional array
Contacts
Insulated leads
(b)
Base
Electrodes
Electrodes
Base
Insulated leads
(a)
Contacts
(c)
Tines
Base
Exposed tip

## Slide 52: Intracellular Recording Electrode
Metal needle with a very fine tip (less than 1.0 µm)
Prepared by electrolytic etching
Metal needle is the anode of an electrolytic cell, and is slowly drawn out of the electrolyte solution (difficult to produce)
Metal must have great strength: stainless steel, platinum-iridium, tungsten, tungsten carbide.
Figure 5.17

## Slide 53: Supported Metal Electrodes
Figure 5.18 (a) Metal-filled glass micropipet. (b) Glass micropipet or probe, coated with metal film.

## Slide 54: Glass Micropipet Electrodes
Figure 5.19 A glass micropipet electrode filled with an electrolytic solution (a) Section of fine-bore glass capillary. (b) Capillary narrowed through heating and stretching. (c) Final structure of glass-pipet microelectrode.

## Slide 55: Microfabricated Microelectrodes
Figure 5.20
Bonding pads
Silicon probe
Exposed
electrodes
Insulated
lead vias
Lead via
Electrode
Silicon probe
Miniature
insulating
chamber
Hole
Si substrate
Exposed tips
SiO2 insulated
Au probes
(a) Beam-lead multiple electrode
(b) Multielectrode silicon probe
(c) Multiple-chamber electrode
Channels
Silicon chip
Contact
metal film
(d) Peripheral-nerve electrode

## Slide 56: Microelectrode Electrical Model
Figure 5.21 (a) Electrode with tip placed within a cell, showing origin of distributed capacitance
N = Nucleus
C = Cytoplasm
Metal rod
Tissue fluid
Membrane
potential
N
C
Insulation
Cell
membrane
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
N = Nucleus
C = Cytoplasm
Shank Capacitance:
Submerged Shaft
Capacitance:
er, e0 = dielectric const.
D = avg. dia. of shank
d = dia. of electrode
t = thickness of
insulation layer
L = length of shank &
shaft, respectively
(5.16)
(5.17)
+
+
+
+
+
+
+
-
-
-
-
-
-
-
Shaft
Submerged
shaft
Shank in
tissue fluid
Shank
inside cell
Electrode tip
inside cell
Tissue
fluid
C

## Slide 57: Microelectrode Electrical Model
(a) Electrode with tip placed within a cell
N = Nucleus
C = Cytoplasm
Metal rod
Tissue fluid
Membrane
potential
N
C
B
A
Reference
electrode
Insulation
Cd
Cell
membrane
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
-
(b) Equivalent circuit
Figure 5.21
B
A
Rmb
Rma
Emb
Ema
Emp
Ri
Re
Cmb
Cma
Cdi
Cd2
Rs
Cw
Cd1
Metal-
electrolyte
interface
Reference
electrode
model
Lead wire
capacitance
Shaft
capacitance
Electrode
resistance
Tissue fluid
resistance
Shank
capacitance
Cell
membrane
Cytoplasm
resistance
To amplifier

## Slide 58: Microelectrode Electrical Model
Emp
Membrane
and
action
potential
Cma
Rma
Cd + Cw
Ema - Emb
E
0
A
B
B
A
Rmb
Rma
Emb
Ema
Emp
Ri
Re
Cmb
Cma
Cdi
Cd2
Rs
Cw
(c) Simplified equivalent circuit
(b) Equivalent circuit
Figure 5.21
Rs, Ri, Re, and Rmb are very
small compared to Rma.

## Slide 59: Glass Micropipet
Figure 5.22 (a) Electrode with its tip placed within a cell, showing the origin of distributed capacitance. (b) Equivalent circuit.
Ema
Rma
Rt
Ri
Re
(b)
Emp
Emb
Rmb
Cmb
Ej
Et
Cma
Cd
A
B
Cell
membrane
Tip
+
-
+
-
+
+
+
-
-
+
-
-
+
-
+
-
+
+
+
+
+
-
-
-
-
-
+
+
+
-
-
-
+
-
-
+
+
+
+
-
-
-
Taper
Internal
electrode
Glass
A
B
To amplifier
Electrolyte
in
micropipet
Stem
(a)
Reference
electrode
Cell membrane
Cytoplasm
N = Nucleus
N
Environmental
fluid
Cd
Rt = electrolyte resistance in shank & tip
Cd = capacitance from micropipet
electrolyte to environmental fluid
Ej = liquid-liquid junction potential
between micropipet electrolyte &
intracellular fluid
Et = tip potential generated by the thin
glass membrane at micropipet tip
Ri = intracellular fluid resistance
Emp = cell membrane potential
Re = extracellular fluid resistance
Internal
electrode
Reference
electrode

## Slide 60: Glass Micropipet
Figure 5.22 (b) Equivalent circuit. (c) Simplified equivalent circuit.
Ema
Rma
Rt
Ri
Re
(b)
Emp
Emb
Rmb
Cmb
Ej
Et
Cma
Cd
A
B
Rt
Em
A
B
Membrane
and
action
potential
(c)
Emp
Em = Ej + Et + Ema- Emb
Cd = Ct
0
Rt = all the series resistance lumped together
(ranges from 1 to 100 MW)
Ct = total distributed capacitance lumped
together (total is tens of pF)
Em = all the dc potentials lumped together
Behaves like a low-pass filter.

## Slide 61: Figure 5.23 (a) Constant-current stimulation (b) Constant-voltage stimulation Charge transfer characteristics of the electrode are very important. Platinum black and Iridium oxide are very good stimulating electrode materials.
(a)
Polarization
potential
Polarization
Polarization
v
v
i
i
t
t
t
t
Ohmic
potential
(b)
Polarization
potential
Stimulating Electrodes

## Slide 62: Practical Hints in Using Electrodes
Ensure that all parts of a metal electrode that will touch the electrolyte are made of the same metal.
  Dissimilar metals have different half-cell potentials making an electrically unstable, noisy junction.
  If the lead wire is a different metal, be sure that it is well insulated.
  Do not let a solder junction touch the electrolyte. If the junction must touch the electrolyte, fabricate the junction by welding or mechanical clamping or crimping.
For differential measurements, use the same material for each electrode.
  If the half-cell potentials are nearly equal, they will cancel and minimize the saturation effects of high-gain, dc coupled amplifiers.
Electrodes attached to the skin frequently fall off.
  Use very flexible lead wires arranged in a manner to minimize the force exerted on the electrode.
  Tape the flexible wire to the skin a short distance from the electrode, making this a stress-relief point.

## Slide 63: Practical Hints in Using Electrodes
A common failure point in the site at which the lead wire is attached to the electrode.
  Repeated flexing can break the wire inside its insulation.
  Prove strain relief by creating a gradual mechanical transition between the wire and the electrode.
  Use a tapered region of insulation that gradually increases in diameter from that of the wire towards that of the electrode as one gets closer and closer to the electrode.
Match the lead-wire insulation to the specific application.
  If the lead wires and their junctions to the electrode are soaked in extracellular fluid or a cleaning solution for long periods of time, water and other solvents can penetrate the polymeric coating and reduce the effective resistance, making the lead wire become part of the electrode.
  Such an electrode captures other signals introducing unwanted noise.
Match your amplifier design to the signal source.
  Be sure that your amplifier circuit has an input impedance that is much greater than the source impedance of the electrodes.